## Slide 1
8. Radiances
Dr. David R. Thompson
Jet Propulsion Laboratory, Imaging Spectroscopy Group

## Slide 2
Goals for this module
Describe factors affecting the physical measurement of radiance spectra
Interpret and apply calibration data to uncalibrated spectral data cubes
Validate radiances in remote field data using in-situ measurements

## Slide 3: Reprise
There are two basic types of imaging spectroscopy data in Earth surface studies.
Transforming from raw instrument values to physical units of radiance is the process of calibration.
Later, we’ll describe the process of estimating surface reflectance from radiance data. This second step is atmospheric correction.
Then, one can use the reflectance spectrum to measure chemical and physical properties of the surface.

## Slide 4: What is the spectral radiance?
Spectral radiance is the radiant flux that is emitted or received by a surface, per unit solid angle per unit projected area per unit wavelength interval.
Per unit wavelength interval
Per unit area
Per unit solid angle
For observations of an extended homogeneous surface, radiance is independent of distance.
For observations of an extended Lambertian surface, radiance is independent of viewing angle.

## Slide 5: Calibration Conventions
Samples (cross-track locations)
Rows (spectral channels)
Spectrometer
Telescope
FPA
Slit
The detector is a two-dimensional focal plane array (FPA). In a typical “pushbroom” design, one is the spatial dimension and the other is the spectral (wavelength) dimension.
To get the second spatial dimension, we ”push” the sensor on a mobile platform.

## Slide 6: Calibration: basic steps
Spectrometer
Telescope
FPA
Slit
Typically, calibration involves a series of transformations and corrections that “work backwards from the focal plane”
Account for electronic effects - the unique, time-dependent radiometric response of each FPA element
Account for optical effects - the spatial and spectral “view” of each FPA element
Tie to the S.I. (absolute spectroradiometry)

## Slide 7: Basic radiometric model
Radiance at FPA
uW nm-1 sr-1 cm-2
FPA Value
DN
Offset
DN
The FPA response is usually linear in illumination. For each crosstrack element x and spectral channel c:
Gain
uW nm-1 sr-1 cm-2 DN-1

## Slide 8: Basic radiometric model
Offset
DN
Gain
uW nm-1 sr-1 cm-2 DN-1
Dark Level
(per element, each flightline)
Radiometric Coefficient
(per channel, each campaign)
Flat field
(per element, each campaign)
This decomposes into different physical effects:
The FPA response is usually linear in illumination. For each crosstrack element x and spectral channel c:

## Slide 9: Components of gain
Estimate RCCs using a NIST-calibrated source
Estimate flat field from a uniform sphere passed across the FOV
Gain is typically separated into two components. A flat field represents the relative response of all elements on the focal plane array. It serves to make their response uniform. Values are typically close to unity.
After this spatial normalization, Radiometric Coefficients (RCCs) map the normalized, offset-subtracted instrument Digital Numbers (DN) to physical radiance units.

## Slide 10: Estimate RCCs using a NIST-calibrated source
Chapman et al., Remote Sensing (2019)
A dark laboratory, or special enclosure, prevents external light from entering.

## Slide 11: Flat field calibration equipment
Chapman et al., Remote Sensing (2019)
Measuring the flat field response image requires a completely uniform illumination from a specially-designed calibration object. Each spectral channel’s response is scaled so that the flat field rows all average to unity.

## Slide 12: Validation method
RTM
images: Mishra et al., Heliyon 2015, wikimedia
reflection
scattering
absorption
We validate flight data by measuring a reflectance signature on the ground, and then converting it to radiance with a Radiative Transfer Model (RTM) that simulates photon propagation through the atmosphere.

## Slide 13: Validation
AVIRIS-NG remote measurement
Model from In-situ measurement
Residual
AVIRIS-NG
MODTRAN simulation
Field reflectance measurements are typically performed at an open playa site to limit scattered light from local vegetation or buildings. The surface should be uniform and flat, and imaged simultaneously with the overflight for similar solar geometry.
Desalpar Playa, India

## Slide 15: Exercise 7: Radiometric Calibration / Validation
Objective: Calibrate digital numbers and “close the loop” with predicted radiances from surface measurements acquired coincidently with an overflight of Ivanpah Playa, NM. Compare alternative Radiative Transfer Models (RTMs). Which is best? Files include:
predicted_radiance_atrem - a prediction of at-sensor radiance following Thompson et al., (RSE 2015).
predicted_radiance_modtran - via the MODTRAN RTM
flat - a flat field frame
rcc.txt - a file of radiometric calibration coefficients
dark - a dark current frame
dn - an imaging spectrometer file containing raw spectrometer data
loc - 3 bands, latitude, longitude, altitude, of the non-orthorectified FN file.
Reflectance was measured between two blue tarps placed on the playa surface, and used in the radiative transfer model. South tarp: 35.60285N, 115.40388W. North tarp: 35.60466, -115.40266